## Slide 1: P2HHV32 96CH LV Aging Test Report
2026.06.15
SZ ENG

## Slide 2
Executive Summary: 4 Units Completed Aging and Passed Available CV Criteria
Completion
4 / 4 Finish
Highest Aging Point
392% / 1.02M
Max CV / Worst
0.825 / 2.88
Key Findings
The detailed lifetime table marks all four requested pipettes as Finish.
All four units reached 392% / 1,020,000 cycles with no total fail, failed channels, or fail analysis recorded.
All available 1 uL CV and worst-channel CV results are below QC/customer limits.

## Slide 3
Optical CV Trend: Maximum 0.825%, Well Below 4% QC Limit
Chart: CV (%)
| 0% | 100% | 200% | 315% | 392%
0401A03 | 0.446 | 0.491 | 0.56 | 0.566 | 0.494
0408A02 | 0.421 | 0.492 | 0.416 | 0.436 | 0.456
0410A07 | 0.382 | 0.386 | 0.718 | 0.724 | 0.415
0411A03 | 0.815 | 0.551 | 0.469 | 0.825 | 0.718
QC Limit
4.0%
Customer Limit
5.0%
Observation
The highest value is 0.825% on 0411A03 at 315%. No trend approaches specification.

## Slide 4
Worst-Channel CV: Maximum 2.88%, Below 5.6% QC Limit
Chart: CV (%)
| 0401A03 | 0408A02 | 0410A07 | 0411A03
Max worst-channel CV | 1.01 | 1.41 | 1.6 | 2.88
QC limit 5.6% | 5.6 | 5.6 | 5.6 | 5.6
Highest worst-channel
2.88%
From 0411A03 at 0%; still within specification.
Decision
All units are below the 5.6% QC limit and 7% customer limit.

## Slide 5
Stage Evidence: Records Complete; Key Scope Notes Flagged
Stage | Cycles | Evidence | Record Result
0% | 0 | Optical CV | Valid CV data for all four units; pass
25%-50% | 65k / 100k / 130k | Water / pressure | PASS
75% | 195k | Skipped | Holiday skip; optical sheet #REF! / #N/A
100% | 260k | Optical CV | All four units pass
125%-175% | 325k / 390k / 455k | Image / check | PASS
200% | 520k | Optical CV | All four units pass
238% / 277% | 620k / 720k | Pressure / water | PASS
315% | 820k | Optical CV | All four units pass
354% | 920k | Pressure / water | PASS
392% | 1,020k | Optical CV | Available data pass; plate4/5 #DIV/0!
Finish in detailed lifetime table

## Slide 6
0401A03 Analysis: Completed 392%; CV Criteria Passed
Finish / 392%
Highest point
392% / 1.02M
Max CV (315%)
0.566%
Max worst (200%)
1.01%
Decision
PASS
Chart: CV (%)
| 0% | 100% | 200% | 315% | 392%
CV across all dispenses | 0.446 | 0.491 | 0.56 | 0.566 | 0.494
Worst-channel CV | 0.92 | 0.91 | 1.01 | 0.81 | 0.91
Reference: CV QC 4%; worst-channel CV QC 5.6%. Axis is scaled to unit-level data.
Detailed Analysis
Trend
CV ranged from 0.446% to 0.566%. The 200% and 315% points were slightly higher, but still far below the 4% QC limit.
Data note
At 392%, CV was 0.494% and worst-channel CV was 0.91%. Plate4/plate5 are #DIV/0! and need retest or documented scope.
Decision
No failure record. Optical CV and worst-channel CV are within specification. Unit passes.
Stage | 0% | 100% | 200% | 315% | 392%
CV | 0.446 | 0.491 | 0.560 | 0.566 | 0.494
Worst | 0.92 | 0.91 | 1.01 | 0.81 | 0.91

## Slide 7
0408A02 Analysis: Completed 392%; CV Criteria Passed
Finish / 392%
Highest point
392% / 1.02M
Max CV (100%)
0.492%
Max worst (315%)
1.41%
Decision
PASS
Chart: CV (%)
| 0% | 100% | 200% | 315% | 392%
CV across all dispenses | 0.421 | 0.492 | 0.416 | 0.436 | 0.456
Worst-channel CV | 0.77 | 0.73 | 0.78 | 1.41 | 0.79
Reference: CV QC 4%; worst-channel CV QC 5.6%. Axis is scaled to unit-level data.
Detailed Analysis
Trend
This unit had the lowest overall CV level among the four units. Worst-channel CV rose to 1.41% at 315%, still with ample margin.
Data note
At 392%, CV was 0.456% and worst-channel CV was 0.79%. Plate4/plate5 are #DIV/0! and need retest or documented scope.
Decision
CV remained stable. No failure record. Unit passes.
Stage | 0% | 100% | 200% | 315% | 392%
CV | 0.421 | 0.492 | 0.416 | 0.436 | 0.456
Worst | 0.77 | 0.73 | 0.78 | 1.41 | 0.79

## Slide 8
0410A07 Analysis: Completed 392%; CV Criteria Passed
Finish / 392%
Highest point
392% / 1.02M
Max CV (315%)
0.724%
Max worst (315%)
1.60%
Decision
PASS
Chart: CV (%)
| 0% | 100% | 200% | 315% | 392%
CV across all dispenses | 0.382 | 0.386 | 0.718 | 0.724 | 0.415
Worst-channel CV | 0.75 | 0.68 | 0.68 | 1.6 | 0.7
Reference: CV QC 4%; worst-channel CV QC 5.6%. Axis is scaled to unit-level data.
Detailed Analysis
Trend
CV increased at 200% and 315%, then returned to 0.415% at 392%. No sustained upward drift was observed.
Data note
Worst-channel CV peaked at 1.60% at 315%, below the 5.6% QC limit. 392% plate4/plate5 are #DIV/0!.
Decision
Mid-life CV rose slightly but stayed far from limits and later recovered. Unit passes.
Stage | 0% | 100% | 200% | 315% | 392%
CV | 0.382 | 0.386 | 0.718 | 0.724 | 0.415
Worst | 0.75 | 0.68 | 0.68 | 1.60 | 0.70

## Slide 9
0411A03 Analysis: Completed 392%; CV Criteria Passed
Finish / 392%
Highest point
392% / 1.02M
Max CV (315%)
0.825%
Max worst (0%)
2.88%
Decision
PASS
Chart: CV (%)
| 0% | 100% | 200% | 315% | 392%
CV across all dispenses | 0.815 | 0.551 | 0.469 | 0.825 | 0.718
Worst-channel CV | 2.88 | 0.99 | 0.57 | 1.76 | 2.44
Reference: CV QC 4%; worst-channel CV QC 5.6%. Axis is scaled to unit-level data.
Detailed Analysis
Trend
This unit had the highest CV and worst-channel CV among the four. The 2.88% worst-channel CV at 0% is still below the 5.6% QC limit.
Data note
0% note: B4/F2/F6 high D% trial 2/3/4. Later available CV results are in spec. 392% plate4/plate5 are #DIV/0!.
Decision
Highest-risk unit in the group, but available CV data did not exceed specification and no failure trend was recorded. Unit passes.
Stage | 0% | 100% | 200% | 315% | 392%
CV | 0.815 | 0.551 | 0.469 | 0.825 | 0.718
Worst | 2.88 | 0.99 | 0.57 | 1.76 | 2.44