## Slide 1: P2HHV32 96CH LV Mold Seal Aging Test Report
2026.06.15
SZ ENG

## Slide 2
Executive Summary: All 5 Mold Seal Pipettes Completed; PASS Recommended
Total units
5
Finish
5
Stop / Fail
0
Max CV
0.86%
Overall conclusion
All 5 units are marked Finish in the RD Lifetime detail sheet and reached 392% / 1,020,000 cycles.
All available 5 uL CV results are below the 4% QC and 5% customer limits; the maximum CV is 0.86% at 100% for 0627A05.
The maximum worst-channel CV is 3.04% at 100% for 0627A05, below the 5.6% QC and 7% customer limits.

## Slide 3
Status Check: Detail Sheet Shows Finish; Summary Sheet Still Needs Sync
SN | Aging summary | RD detail | Highest point | Report judgment
P2HHV3220250407A02 | 200% running | Finish | 392% / 1.02M | PASS
P2HHV3220250627A05 | 200% running | Finish | 392% / 1.02M | PASS
P2HHV3220250701A05 | 200% running | Finish | 392% / 1.02M | PASS
P2HHV3220250627A01 | 200% running | Finish | 392% / 1.02M | PASS
P2HHV3220250702A04 | 200% running | Finish | 392% / 1.02M | PASS
Sync gap
All 5 units are Finish in the RD detail sheet, while the aging summary still shows 200% (520000) and 200%-238% running.
Judgment basis
This report uses the RD Lifetime detail sheet, optical CV sheet top metrics, and staged PASS records as the judgment basis.

## Slide 4
5 uL Optical CV: Maximum 0.86%, Well Below the 4% QC Limit
Chart: CV (%)
| 0407A02 | 0627A05 | 0701A05 | 0627A01 | 0702A04
Max CV | 0.54 | 0.86 | 0.69 | 0.72 | 0.64
QC 4% | 4 | 4 | 4 | 4 | 4
Customer 5% | 5 | 5 | 5 | 5 | 5
Highest CV
0.86%
0627A05 / 100%
QC limit
4.0%
Observation
Every unit's maximum CV is below 1%, leaving strong specification margin.

## Slide 5
Worst-Channel CV: Maximum 3.04%, Below the 5.6% QC Limit
Chart: CV (%)
| 0407A02 | 0627A05 | 0701A05 | 0627A01 | 0702A04
Max worst-channel CV | 3 | 3.04 | 2.57 | 1.15 | 1.92
QC 5.6% | 5.6 | 5.6 | 5.6 | 5.6 | 5.6
Highest value
3.04%
0627A05 / 100%, with a C8 LOW %D note.
Conclusion
All units' worst-channel CV values are below the 5.6% QC and 7% customer limits.

## Slide 6
0407A02 Unit Summary: PASS
Finish / 392%
Highest point
392% / 1.02M
Max CV (100% / 315%)
0.54%
Max worst ch. (315%)
3.00%
Judgment
PASS
Chart: CV (%)
| 0% | 100% | 200% | 315% | 392%
Overall CV | 0.44 | 0.54 | 0.42 | 0.54 | 0.35
Worst-channel CV | 0.76 | 1.01 | 0.73 | 3 | 0.71
Spec reference: 5 uL CV QC 4% / customer 5%; worst-channel CV QC 5.6% / customer 7%.
Unit summary
Trend
CV remained between 0.35% and 0.54% across all stages, well below the 4% QC limit; worst-channel CV peaked at 3.00% at 315%, still below the 5.6% QC limit.
Data note
392% plate4/plate5 are #DIV/0!.
Conclusion
The detailed sheet shows Finish at 392% / 1,020,000 cycles. Current CV results meet the specification; recommend PASS for archive.
Stage | 0% | 100% | 200% | 315% | 392%
CV | 0.44 | 0.54 | 0.42 | 0.54 | 0.35
Worst ch. | 0.76 | 1.01 | 0.73 | 3.00 | 0.71

## Slide 7
0627A05 Unit Summary: PASS
Finish / 392%
Highest point
392% / 1.02M
Max CV (100%)
0.86%
Max worst ch. (100%)
3.04%
Judgment
PASS
Chart: CV (%)
| 0% | 100% | 200% | 315% | 392%
Overall CV | 0.36 | 0.86 | 0.53 | 0.36 | 0.56
Worst-channel CV | 0.67 | 3.04 | 0.76 | 0.99 | 0.8
Spec reference: 5 uL CV QC 4% / customer 5%; worst-channel CV QC 5.6% / customer 7%.
Unit summary
Trend
CV and worst-channel CV peaked at 100%, but 0.86% CV and 3.04% worst-channel CV remain below the limits.
Data note
100% has a C8 LOW %D in plate1 note; 315% notes the new validation with only 3 plates tested; 315%/392% plate4/plate5 are #DIV/0!.
Conclusion
The detailed sheet shows Finish and current CV results meet the specification. Recommend PASS, retaining the C8 D% and 3-plate scope notes.
Stage | 0% | 100% | 200% | 315% | 392%
CV | 0.36 | 0.86 | 0.53 | 0.36 | 0.56
Worst ch. | 0.67 | 3.04 | 0.76 | 0.99 | 0.80

## Slide 8
0701A05 Unit Summary: PASS
Finish / 392%
Highest point
392% / 1.02M
Max CV (100%)
0.69%
Max worst ch. (100%)
2.57%
Judgment
PASS
Chart: CV (%)
| 0% | 100% | 200% | 315% | 392%
Overall CV | 0.42 | 0.69 | 0.47 | 0.41 | 0.37
Worst-channel CV | 0.62 | 2.57 | 0.87 | 0.68 | 0.72
Spec reference: 5 uL CV QC 4% / customer 5%; worst-channel CV QC 5.6% / customer 7%.
Unit summary
Trend
A temporary increase was observed at 100%; values then returned lower and stayed stable from 200% to 392%. Maximum CV was 0.69%.
Data note
392% plate4/plate5 are #DIV/0!.
Conclusion
The detailed sheet shows Finish. Current CV and worst-channel CV results are within specification; recommend PASS.
Stage | 0% | 100% | 200% | 315% | 392%
CV | 0.42 | 0.69 | 0.47 | 0.41 | 0.37
Worst ch. | 0.62 | 2.57 | 0.87 | 0.68 | 0.72

## Slide 9
0627A01 Unit Summary: PASS
Finish / 392%
Highest point
392% / 1.02M
Max CV (100%)
0.72%
Max worst ch. (100%)
1.15%
Judgment
PASS
Chart: CV (%)
| 0% | 100% | 200% | 315% | 392%
Overall CV | 0.49 | 0.72 | 0.5 | 0.35 | 0.38
Worst-channel CV | 0.81 | 1.15 | 0.63 | 0.66 | 0.54
Spec reference: 5 uL CV QC 4% / customer 5%; worst-channel CV QC 5.6% / customer 7%.
Unit summary
Trend
CV peaked at 0.72% at 100%, then decreased to 0.35%-0.50%; maximum worst-channel CV was 1.15%, leaving strong margin.
Data note
315% notes the new validation with only 3 plates tested; 315%/392% plate4/plate5 are #DIV/0!.
Conclusion
The detailed sheet shows Finish. Current CV data are stable and meet the specification; recommend PASS.
Stage | 0% | 100% | 200% | 315% | 392%
CV | 0.49 | 0.72 | 0.50 | 0.35 | 0.38
Worst ch. | 0.81 | 1.15 | 0.63 | 0.66 | 0.54

## Slide 10
0702A04 Unit Summary: PASS
Finish / 392%
Highest point
392% / 1.02M
Max CV (100%)
0.64%
Max worst ch. (315%)
1.92%
Judgment
PASS
Chart: CV (%)
| 0% | 100% | 200% | 315% | 392%
Overall CV | 0.47 | 0.64 | 0.45 | 0.44 | 0.37
Worst-channel CV | 0.72 | 0.82 | 0.78 | 1.92 | 0.61
Spec reference: 5 uL CV QC 4% / customer 5%; worst-channel CV QC 5.6% / customer 7%.
Unit summary
Trend
CV peaked at 0.64% and then decreased to 0.37%; worst-channel CV was 1.92% at 315%, still well below the 5.6% QC limit.
Data note
75% is skipped/#REF; 315% notes the new validation with only 3 plates tested; 315%/392% plate4/plate5 are #DIV/0!.
Conclusion
The detailed sheet shows Finish and current CV metrics meet the specification; recommend PASS.
Stage | 0% | 100% | 200% | 315% | 392%
CV | 0.47 | 0.64 | 0.45 | 0.44 | 0.37
Worst ch. | 0.72 | 0.82 | 0.78 | 1.92 | 0.61