## Slide 1: P2HHV32 96CH LV Rod Seal Aging Test Report
2026.06.15
SZ ENG

## Slide 2
Executive Summary: Aging Test Conclusions for 5 Rod Seal Pipettes
Total Units
5
Finish
3
Fail / Stop
2
QC Risk
1
Overall Conclusion
0409A03 and 0409A02 completed 392% / 1,020,000 cycles; available CV and worst-channel CV meet QC/customer specifications.
0627A02 reached 6.83% CV at 175% optical verification and had D3/G9 aspiration issues; classify as stopped / failed.
0627A03 had aspiration failures at 50%/175%; 0701A01 finished but 392% CV reached 4.155%-4.16%, above the 4% QC limit.

## Slide 3
Status Check: Detailed Table and Summary Sheet Are Not Fully Synchronized
SN | Summary Sheet | RD Detailed Table | Highest Point | Report Decision
P2HHV3220250627A02 | 175% stopped; rework/FA | STOPED-(150%-175%) | 175% / 455k | FAIL / STOP
P2HHV3220250409A03 | 200% running | Finish | 392% / 1.02M | PASS
P2HHV3220250409A02 | 200% running | Finish | 392% / 1.02M | PASS
P2HHV3220250627A03 | 150%-175% stopped | stopped-(150%-175%) | 175% / 455k | FAIL / STOP
P2HHV3220250701A01 | 200% running | Finish | 392% / 1.02M | QC RISK
Synchronization Gap
0409A03, 0409A02, and 0701A01 are Finish in the detailed table, while the summary sheet still shows 200% running.
Decision Basis
This report uses the RD Lifetime detailed table, optical CV sheets, and explicit fail notes as the decision basis.

## Slide 4
5 uL Optical CV: Two Stopped/Risk Units Need Priority Handling
Chart: CV (%)
| 0627A02 | 0409A03 | 0409A02 | 0627A03 | 0701A01
Max CV | 6.83 | 0.52 | 0.74 | 0.47 | 4.16
QC 4% | 4 | 4 | 4 | 4 | 4
Customer 5% | 5 | 5 | 5 | 5 | 5
Highest CV
6.83%
0627A02 / 175%
QC Risk
4.16%
0701A01 / 392%
Note
0409A03/0409A02 CV stayed far below the 4% QC limit.

## Slide 5
Worst-Channel CV: Maximum 3.45%, Below the 5.6% QC Limit
Chart: CV (%)
| 0627A02 | 0409A03 | 0409A02 | 0627A03 | 0701A01
Max worst-channel CV | 0.81 | 1.07 | 3.45 | 0.66 | 2.3
QC 5.6% | 5.6 | 5.6 | 5.6 | 5.6 | 5.6
Highest Value
3.45%
0409A02 / 200%, with B5 LOW %D note.
Conclusion
Worst-channel CV did not exceed 5.6% QC. Failed units are driven mainly by aspiration/leak and overall CV issues.

## Slide 6
Failure and Risk Evidence: 2 Stopped Units, 1 Finish Unit Requiring Review
SN | Stage | Evidence / Note | Impact | Recommendation
0627A02 | 175% | CV 6.83%; D3/G9 aspiration issue | Above QC/customer; stopped | 
0627A03 | 50%-175% | E10 aspiration fail; CH19/CH58 high leak; B7 aspiration FAIL | Process failure; stopped | 
0701A01 | 392% | CV 4.155%-4.16%; H4 LOW %D; leak 23.46 pa/s | Above 4% QC; below 5% customer | 
0409A02 | 200% | B5 LOW %D; worst-channel CV 3.45% | CV in spec; single-point risk | 
0409A03 | All | No failure record; max CV 0.52% | Stable pass | 

## Slide 7
0627A02 Unit Summary: FAIL / STOP
STOPED-(150%-175%)
Highest Point
175% / 455k
Max CV (175%)
6.83%
Max Worst (100%)
0.81%
Decision
FAIL / STOP
Chart: CV (%)
| 0% | 100% | 175%
Overall CV | 0.46 | 0.65 | 6.83
Worst-channel CV | 0.78 | 0.81 | 0.77
Reference: 5 uL CV QC 4% / customer 5%; worst-channel CV QC 5.6% / customer 7%.
Unit Summary
Trend
CV reached 6.83% at 175% optical verification, above both the 4% QC limit and 5% customer limit.
Data Note
Main table notes D3 low/failed aspiration; optical verification records D3/G9 aspiration issues. Summary indicates stopped test with rework/disassembly analysis.
Decision
Fail/stop. Do not archive this unit as aging-test pass.
Stage | 0% | 100% | 175%
CV | 0.46 | 0.65 | 6.83
Worst | 0.78 | 0.81 | 0.77

## Slide 8
0409A03 Unit Summary: PASS
Finish
Highest Point
392% / 1.02M
Max CV (200% / 315%)
0.52%
Max Worst (200%)
1.07%
Decision
PASS
Chart: CV (%)
| 0% | 100% | 200% | 315% | 392%
Overall CV | 0.43 | 0.51 | 0.52 | 0.52 | 0.41
Worst-channel CV | 0.7 | 0.94 | 1.07 | 0.73 | 0.64
Reference: 5 uL CV QC 4% / customer 5%; worst-channel CV QC 5.6% / customer 7%.
Unit Summary
Trend
CV stayed stable at 0.41%-0.52%, far below the 4% QC limit.
Data Note
At 392%, plate4/plate5 are #DIV/0!. Retest or document that the final optical summary uses the first three plates only.
Decision
Detailed table is Finish and available CV / worst-channel CV are within specification. Recommend pass.
Stage | 0% | 100% | 200% | 315% | 392%
CV | 0.43 | 0.51 | 0.52 | 0.52 | 0.41
Worst | 0.70 | 0.94 | 1.07 | 0.73 | 0.64

## Slide 9
0409A02 Unit Summary: PASS
Finish
Highest Point
392% / 1.02M
Max CV (200%)
0.74%
Max Worst (200%)
3.45%
Decision
PASS
Chart: CV (%)
| 0% | 100% | 200% | 315% | 392%
Overall CV | 0.47 | 0.58 | 0.74 | 0.42 | 0.43
Worst-channel CV | 0.73 | 0.91 | 3.45 | 0.65 | 0.75
Reference: 5 uL CV QC 4% / customer 5%; worst-channel CV QC 5.6% / customer 7%.
Unit Summary
Trend
Maximum CV was 0.74%, well below the 4% QC limit; CV recovered to 0.42%-0.43% at 315%-392%.
Data Note
200% note: B5 LOW %D in plate3. Worst-channel CV peaked at 3.45%, still below the 5.6% QC limit.
Decision
Completed 392% and CV criteria passed. Recommend pass, with B5 D% note retained.
Stage | 0% | 100% | 200% | 315% | 392%
CV | 0.47 | 0.58 | 0.74 | 0.42 | 0.43
Worst | 0.73 | 0.91 | 3.45 | 0.65 | 0.75

## Slide 10
0627A03 Unit Summary: FAIL / STOP
stopped-(150%-175%)
Highest Point
175% / 455k
Max CV (100%)
0.47%
Max Worst (0%)
0.66%
Decision
FAIL / STOP
Chart: CV (%)
| 0% | 100%
Overall CV | 0.43 | 0.47
Worst-channel CV | 0.66 | 0.61
Reference: 5 uL CV QC 4% / customer 5%; worst-channel CV QC 5.6% / customer 7%.
Unit Summary
Trend
Available optical CV was normal at 0%-100%, but the lifetime process recorded repeated aspiration/leak issues.
Data Note
50% E10 aspiration failure followed by rework continuation; 150% CH19/CH58 high leak rate; 175% B7 aspiration failure with FAIL record.
Decision
Although available CV did not exceed limits, the explicit 175% FAIL record drives fail/stop decision.
Stage | 0% | 100%
CV | 0.43 | 0.47
Worst | 0.66 | 0.61

## Slide 11
0701A01 Unit Summary: QC RISK
Finish
Highest Point
392% / 1.02M
Max CV (392%)
4.16%
Max Worst (315%)
2.30%
Decision
QC RISK
Chart: CV (%)
| 0% | 100% | 200% | 315% | 392%
Overall CV | 0.4 | 0.74 | 1.08 | 1.4 | 4.16
Worst-channel CV | 0.73 | 1.73 | 0.78 | 2.3 | 0.93
Reference: 5 uL CV QC 4% / customer 5%; worst-channel CV QC 5.6% / customer 7%.
Unit Summary
Trend
CV increased with aging and reached 4.155%-4.16% at 392%, above the 4% QC limit but below the 5% customer limit.
Data Note
200%/354% notes: H4 LOW %D. At 354%, pressure fixture leak rate was recorded as 23.46 pa/s for 200 uL.
Decision
Detailed table is Finish, but retest/review the 392% optical result and H4-related findings before archival.
Stage | 0% | 100% | 200% | 315% | 392%
CV | 0.40 | 0.74 | 1.08 | 1.40 | 4.16
Worst | 0.73 | 1.73 | 0.78 | 2.30 | 0.93