## Slide 1: P2HHV32 96CH LV Three-Type Aging Test Comparison Report
2026.06.15
SZ ENG

## Slide 2
Executive Summary: Main Difference Is Concentrated in Rod Seal
Compared types
3
Total units
14
Reached 392%
12
STOP / Risk
3
Key conclusion
Mold Oring Rework and Mold Seal both show 100% Finish / PASS, and both have max CV below 1%, leaving strong margin.
Rod Seal has only 2/5 clear PASS units, plus 2 STOP/FAIL and 1 QC risk unit; it is the only type not recommended for direct release.
Rod Seal issues are not driven by broad worst-channel CV exceedance, but by low aspiration, leakage/leak-rate notes, and 392% CV risk.
Direction
1. Archive Mold Oring Rework and Mold Seal as PASS.
2. Hold Rod Seal until failure RCA, structural/assembly review, and retest are complete.
3. Sync the aging summary before formal archive.
4. Align the retest or written explanation for 392% plate4/plate5 #DIV/0!.

## Slide 3
Three-Type Status Snapshot: Rod Seal Has STOP/FAIL and QC Risk
Type | Units | Finish | PASS | Risk | Fail/Stop | Max CV | Max worst ch. | Overall judgment
Mold Oring Rework | 4 | 4 | 4 | 0 | 0 | 0.825% | 2.88% | PASS
Rod Seal | 5 | 3 | 2 | 1 | 2 | 6.83% | 3.45% | HOLD / RCA
Mold Seal | 5 | 5 | 5 | 0 | 0 | 0.86% | 3.04% | PASS
Mold Oring
4/4 reached 392% / 1.02M
Final PASS: 4/4; Fail/Stop: 0/4
Judgment: PASS
Rod Seal
3/5 reached 392%; 2/5 stopped around 175%
Final PASS: 2/5; Fail/Stop: 2/5
Judgment: HOLD / RCA
Mold Seal
5/5 reached 392% / 1.02M
Final PASS: 5/5; Fail/Stop: 0/5
Judgment: PASS

## Slide 4
Completion and Final PASS Rate: Rod Seal Lags Clearly
Chart: Rate (%)
| Mold Oring | Rod Seal | Mold Seal
Finish rate | 100 | 60 | 100
Final PASS rate | 100 | 40 | 100
Best PASS rate
100%
Mold Oring / Mold Seal
Rod Seal PASS
40%
Risk
Rod Seal has 2 STOP/FAIL units and 1 QC risk unit.

## Slide 5
Maximum Optical CV: Rod Seal Exceeds QC Limit at 175%
Chart: CV (%)
| Mold Oring | Rod Seal | Mold Seal
Max CV | 0.825 | 6.83 | 0.86
QC 4% | 4 | 4 | 4
Customer 5% | 5 | 5 | 5
Highest value
6.83%
Rod Seal / 0627A02 / 175%, with D3/G9 low aspiration.
Passing types
Mold Oring and Mold Seal both have max CV below 1%.

## Slide 6
Maximum Worst-Channel CV: All Three Types Below the 5.6% QC Limit
Chart: CV (%)
| Mold Oring | Rod Seal | Mold Seal
Max worst-channel CV | 2.88 | 3.45 | 3.04
QC 5.6% | 5.6 | 5.6 | 5.6
Customer 7% | 7 | 7 | 7
Highest value
3.45%
Rod Seal / 0409A02 / 200%
QC limit
5.6%
Interpretation
Rod Seal failures lean mechanical/aspiration/leak related, not broad worst-channel CV exceedance.

## Slide 7
Mold Oring Rework: Overall PASS
PASS
Sample count
4
Finish / PASS
4/4 / 4/4
Max CV
0.825%
Max worst ch.
2.88%
Type conclusion
All 4 units are Finish and completed 392%; all available CV and worst-channel CV results are below limits, giving the best margin among the three types.
The summary sheet still shows 238% running; 392% plate4/plate5 are #DIV/0!; 0411A03 has an initial 0% high %D note, but later CV remains within spec.
SN | Status | Judgment | Max CV | Max worst | Note
0401A03 | Finish | PASS | 0.566% | 1.01% | 392% plate4/plate5 #DIV/0!
0408A02 | Finish | PASS | 0.492% | 1.41% | 392% plate4/plate5 #DIV/0!
0410A07 | Finish | PASS | 0.724% | 1.60% | 392% plate4/plate5 #DIV/0!
0411A03 | Finish | PASS | 0.825% | 2.88% | 0% high %D; later CV within spec

## Slide 8
Rod Seal: HOLD and RCA Required
HOLD / RCA
Sample count
5
Finish / PASS
3/5 / 2/5
Max CV
6.83%
Max worst ch.
3.45%
Type conclusion
This is the only type with STOP/FAIL and QC risk records; issues concentrate in low aspiration, leakage/leak-rate notes, and 392% CV QC risk.
0627A02 has 6.83% CV at 175% with D3/G9 low aspiration; 0627A03 failed between 150%-175%; 0701A01 is Finish but has 4.16% CV at 392%, judged QC risk.
SN | Status | Judgment | Max CV | Max worst | Note
0627A02 | STOP 150%-175% | FAIL/STOP | 6.83% | 0.81% | D3/G9 low aspiration; rework/teardown
0409A03 | Finish | PASS | 0.52% | 1.07% | 392% plate4/plate5 #DIV/0!
0409A02 | Finish | PASS | 0.74% | 3.45% | 200% B5 LOW %D
0627A03 | STOP 150%-175% | FAIL/STOP | 0.47% | 0.66% | E10/B7 no aspiration; CH19/58 leak
0701A01 | Finish | QC RISK | 4.16% | 2.30% | H4 LOW %D; leak rate 23.46 pa/s

## Slide 9
Mold Seal: Overall PASS
PASS
Sample count
5
Finish / PASS
5/5 / 5/5
Max CV
0.86%
Max worst ch.
3.04%
Type conclusion
All 5 units are Finish and completed 392%; max CV is 0.86% and max worst-channel CV is 3.04%, both within specification.
The summary sheet still shows 200% running; 392% plate4/plate5 are #DIV/0!; some 315% records used only 3 plates under the new validation.
SN | Status | Judgment | Max CV | Max worst | Note
0407A02 | Finish | PASS | 0.54% | 3.00% | 392% plate4/plate5 #DIV/0!
0627A05 | Finish | PASS | 0.86% | 3.04% | 100% C8 LOW %D; 3-plate scope
0701A05 | Finish | PASS | 0.69% | 2.57% | 392% plate4/plate5 #DIV/0!
0627A01 | Finish | PASS | 0.72% | 1.15% | 315% new validation: 3 plates
0702A04 | Finish | PASS | 0.64% | 1.92% | 75% skip/#REF; 3-plate scope

## Slide 10
Failure Mode and Data Quality Comparison
1
Mold Oring Rework
Failure / risk
No STOP/FAIL; only 0411A03 has an early 0% high %D note, and later CV stays within spec.
Data quality
Summary sheet lags at 238%; 392% plate4/plate5 are #DIV/0!.
2
Rod Seal
Failure / risk
2 STOP/FAIL units: low/no aspiration and high leak-rate notes; 1 Finish unit has 392% CV 4.16% QC risk.
Data quality
Some Finish units still show 200% running in the summary; 392% plate4/plate5 are #DIV/0!.
3
Mold Seal
Failure / risk
No STOP/FAIL; 0627A05 has a C8 LOW %D note, but CV / worst-channel CV remain within spec.
Data quality
Summary sheet lags at 200%; 392% plate4/plate5 are #DIV/0!; some 315% records use only 3 plates.
Common action
All three types have aging-summary vs. RD-detail sync gaps and 392% plate4/plate5 #DIV/0! archive-scope questions. Align retest or written explanation before final archive.

## Slide 11
Final Recommendation: Two Types Pass; Rod Seal Requires RCA Before Release
1
Recommend PASS
Mold Oring Rework: 4/4 Finish, 4/4 PASS, max CV 0.825%, max worst-channel CV 2.88%; best optical margin among the three types.
2
Recommend PASS
Mold Seal: 5/5 Finish, 5/5 PASS, max CV 0.86%, max worst-channel CV 3.04%; archive in the PASS direction.
3
Hold release
Rod Seal: 2/5 STOP/FAIL and 1/5 QC risk. Complete failed-unit teardown, aspiration/leak RCA, structure/assembly review, and retest before final judgment.
Unified archive actions
Sync the aging summary; align retest or written explanation for 392% plate4/plate5 #DIV/0!; retain all LOW %D, low aspiration, leak-rate, and 3-plate validation notes.