## Slide 1: P2HHV32 96CH LV Two O-ring Type Aging Test Comparison Report
2026.06.16
SZ ENG

## Slide 2
Executive Summary: Mold Oring Completed Well; OTP-O Ring Risk Is Concentrated
O-ring types
Mold/OTP O
ring
Total samples
18 units
Reached 392%/417%
16 units
Risk / Fail / Hold
5 units
Key Findings
Mold Oring includes 13 units; all completed 392% or 417%. Units 0407A01 and 0702A03 require focused follow-up.
OTP-O ring includes 5 units: only 2 clear PASS, 2 STOP/FAIL, and 1 QC risk. Recommend HOLD until RCA is closed.
Failure patterns differ: Mold Oring is mainly local leakage/O-ring rework and individual-channel D% risk; OTP-O ring is mainly low/no aspiration and leakage risk.

## Slide 3
Two O-ring Type Test Status Snapshot
Type | Sample | Complete | PASS | QC risk | HOLD | STOP/FAIL | Max CV | Max worst ch. | Overall decision
Mold Oring | 13 | 13 | 11 | 2 | 0 | 0 | 6.6% | 55.15% | PASS with Risk Notes
OTP-O ring | 5 | 3 | 2 | 1 | 0 | 2 | 6.83% | 3.45% | HOLD / RCA
Mold Oring
Complete: 13/13
PASS: 11/13; Risk/Hold/Fail: 2/13
Decision: PASS with Risk Notes
OTP-O ring
Complete: 3/5
PASS: 2/5; Risk/Hold/Fail: 3/5
Decision: HOLD / RCA

## Slide 4
Completion and PASS Rates: OTP-O Ring Is Clearly Lower Than Mold Oring
Chart: Rate (%)
| Mold Oring | OTP-O ring
Completion rate | 100 | 60
PASS rate | 85 | 40
Risk/Fail/Hold rate | 15 | 60
Mold Oring complete
100%
OTP PASS
40%
Risk Signal
OTP-O ring has 3/5 risk or failed units; Mold Oring has 2/13 units requiring follow-up.

## Slide 5
CV across all dispenses:
Chart: CV (%)
| Mold Oring | OTP-O ring
Maximum CV | 6.6 | 6.83
QC 4% | 4 | 4
Customer 5% | 5 | 5
Highest
6.83%
OTP-O ring / 0627A02 / 175%, D3/G9 low aspiration.
Mold Oring
0407A01 reached 6.60% CV at 379%, related to D2 leakage / O-ring replacement.

## Slide 6
CV of the worst channel: Mold Oring Has a Significant Individual Spike
Chart: CV (%)
| Mold Oring | OTP-O ring
Max worst-channel CV | 55.15 | 3.45
QC 5.6% | 5.6 | 5.6
Customer 7% | 7 | 7
Highest
55.15%
Mold Oring / 0407A01 / 379%, D2 leakage.
Interpretation
OTP-O ring max worst-channel CV is 3.45%, below 5.6% QC; its main risks are aspiration and leakage.

## Slide 7
Mold Oring Type Analysis: Overall Can Move Forward, with Risk Closure Required
PASS with risk notes
Sample size
13 units
Complete / PASS
13/13 / 11/13
Max CV
6.60%
Max worst ch.
55.15%
Type Conclusion
All 13 units included in this analysis completed aging to 392% or 417%; 11 units are recommended as PASS, while 2 units retain QC risk / rework-related risk.
Risks are concentrated in P2HHV3120250407A01 with D2 leakage / O-ring replacement and P2HHV3120250702A03 with H3 low %D plus a worst-channel CV spike.
Status | Qty | Notes
PASS | 11 | Most units completed 392%/417% aging with CV within specification.
QC risk | 2 | D2 leakage / H3 low %D and worst-channel CV excursions.
HOLD/RUNNING | 0 | No HOLD/RUNNING units in the included sample set.

## Slide 8
OTP-O Ring Type Analysis: HOLD and RCA Recommended
HOLD / RCA
Sample size
5 units
Complete / PASS
3/5 / 2/5
Max CV
6.83%
Max worst ch.
3.45%
Type Conclusion
The sample size is small, but the risk ratio is the highest: 2 of 5 units are STOP/FAIL, 1 unit is QC risk, and only 2 units are clear PASS.
Failures are concentrated in low/no aspiration, abnormal leakage rate, and 392% CV risk; 0627A02 and 0627A03 stopped, while 0701A01 finished but reached 4.16% CV at 392%.
SN | Status | Decision | Max CV | Max worst | Notes
P2HHV3220250627A02 | STOP 150%-175% | FAIL/STOP | 6.83% | 0.81% | D3/G9 low aspiration at 175% optical check; teardown/rework.
P2HHV3220250627A03 | STOP 150%-175% | FAIL/STOP | 0.47% | 0.66% | E10/B7 no aspiration; CH19/CH58 high leakage rate.
P2HHV3220250701A01 | Finish / 392% | QC RISK | 4.16% | 2.30% | H4 LOW %D; 200 uL leakage rate 23.46 pa/s.

## Slide 9
Final Recommendation: Conditional Mold Oring Release; Hold OTP-O Ring
1
Conditional PASS
Mold Oring: Move the 11 normal completed units toward PASS; close 0407A01 and 0702A03 as separate risk items before archiving.
2
HOLD / RCA
OTP-O ring: 2/5 STOP/FAIL and 1/5 QC risk. Complete failed-part teardown, aspiration/leakage RCA, structure/assembly review, and retest before final decision.