## Slide 1
96CH LV Aging Test Workload and Result Overview
opentrons
Types covered
2
Units tested
19
Reached 1,020,000/1,085,000
16
PASS / Overall rate
15 / 79%
Fail closed
4
Work Completed
Re-matched 19 listed units and summarized them under two groups: Mold Oring and OTP-O ring.
Reviewed RD detail status, maximum aging node, CV, worst-channel CV, LOW %D, aspiration, no-aspiration, and leakage notes.
Coverage spans 0 to 1,020,000/1,085,000 cycles; 16 units reached endpoint and 15 units are counted as PASS.
Core workload: 19 units x multi-node validation
Type 1: Mold Oring
14 units: all counted as Mold Oring
PASS with risk closure
Units
14
Endpoint
13/14
PASS
13/14
Pass rate
93%
Risk/Fail
1
Completion
93%
PASS rate
93%
Max coverage: 1,085,000 cycles; main endpoint: 1,020,000 cycles
Workload: 14 units included; 13 units reached 1,020,000 or 1,085,000 cycles.
Pass rate: 13/14 PASS; 0702A03 and 0410A04 are counted as passed.
Risk review: 0407A01 D2 leakage - 1
Type 2: OTP-O ring
5 units: ROD SEAL PIPETTE (10433+10425+10026OTP)
FAIL closed / HOLD
Units
5
Endpoint
3/5
PASS
2/5
Pass rate
40%
Risk/Fail
3
Completion
60%
PASS rate
40%
3 units reached 1,020,000 cycles; 2 stopped at 150%-175%
Workload: 5 units included; coverage from 0 to 1,020,000 cycles, with 2 STOP/FAIL units.
Pass rate: 2/5 PASS; 1 QC risk; 2 units have analysis reports and closed Fail disposition.
Main risks: D3/G9 low aspiration, E10/B7 no aspiration, leakage, and 4.16% CV at 1,020,000 cycles.
Conclusion
Mold Oring aging validation shows a clear lifetime improvement: compared with the original 270,000-cycle design life, current units passed 1,020,000-1,085,000 cycles, reaching about 3.8-4.0x life (approximately 278%-302% improvement).
Note: Some P2HHV32 rows in the aging summary table lag behind; final statistics use the latest RD detail nodes.