## Slide 1
P1000H O-ring
Data Analysis Report
Data-focused analysis of aspiration behavior and optical CV
1,020,000 cycles / 392.3% life
Report Type
Data Analysis
Trends / Anomaly Nodes
Plate-level Scope
Unit-level Evidence Chain
SZ ENG | 2026-06-04
P1000H O-ring Data Analysis | 2026-06-04 | 1/13

## Slide 2
1. Data Scope and Analysis Method
This report focuses on aspiration / LOW %D and optical CV. Pressure data is treated only as supporting seal-context data, not as the primary judgment basis.
Opentrons
P1000H O-ring Data Analysis | 2026-06-04 | 2/13
Field
Data Scope
Sample Count
5 production P1KH 96-channel P1000H units
Test Type
1000 uL cycles / O-ring-related performance observation
Cycle Records
Each unit has 15 cycle rows; 10 rows contain analyzable numeric CV data
Latest Node
1,020,000 cycles / 392.3% life / Finish x5
Key Data Fields
CV across all dispenses, plate1-5 CV, worst-channel CV, LOW %D / fail / note
Key Scope Change
At the latest 1020k node, only plate1-3 are valid; plate4/5 are #DIV/0!
Analysis principle: treat the O-ring as an engineering hypothesis that may affect aspiration consistency and optical CV; use the data to grade risk, and do not treat inference as a confirmed root cause.

## Slide 3
2. Metric Definitions: Data That Indicates O-ring Risk
The O-ring does not directly generate optical values; it affects aspiration/dispense stability through sealing and friction, which then appears in optical results.
Opentrons
P1000H O-ring Data Analysis | 2026-06-04 | 3/13
Category
Metric
Analytical Meaning
Direct Observation
LOW %D / low aspiration
Low aspiration or local channel aspiration shortage is the closest observable signal to an O-ring sealing issue.
Direct Observation
Optical CV across all dispenses
Overall dispensing consistency; an increase suggests higher channel-to-channel or plate-to-plate variation.
Direct Observation
Worst-channel CV
Used to locate abnormal peaks from a single channel or a small group of channels.
Data Quality
Valid Plate Count
The latest 1020k node has only 3 tested plates, so it is not fully equivalent to historical 5-plate data.
Engineering Inference
O-ring Risk
O-ring risk increases when LOW %D, CV peaks, and channel-localized evidence appear together.
• This report prioritizes data patterns: overall elevation, localized channel anomalies, and whether LOW %D / fail notes are present.
• O-ring conclusions must be confirmed through teardown or before/after replacement retesting; this report grades the strength of data support.

## Slide 4
3. CV Trend: Different Units Show Different Variation Patterns
A05 is a late-stage spike; A01 was high in the early/mid stage and then improved; A04 is elevated during 195k-520k and again at the latest node.
Opentrons
P1000H O-ring Data Analysis | 2026-06-04 | 4/13
CV across all dispenses by cycle
0
1
2
3
4
5
6
7
8
9
0
260k
520k
820k
1020k
CV 5 reference
0624A04
1217A05
0311A04
0624A01
0624A05
Unit
Latest CV
Max CV point
N
0624A04
4.379
520k / 5.642
10
1217A05
1.487
520k / 6.195
10
0311A04
0.677
0 / 1.957
10
0624A01
1.633
520k / 8.406
10
0624A05
6.839
1020k / 6.839
10

## Slide 5
4. Worst-channel CV: Channel-level Peak Localization
Worst-channel CV helps identify local sealing or aspiration instability, rather than relying only on overall CV.
Opentrons
P1000H O-ring Data Analysis | 2026-06-04 | 5/13
Worst-channel CV by cycle
0
2
4
6
8
10
12
14
0
260k
520k
820k
1020k
0624A04
1217A05
0311A04
0624A01
0624A05
Unit
Max worst-channel
Latest worst-channel
Risk
0624A04
520k / 11.204
1020k / 1.585
High
1217A05
520k / 7.419
1020k / 2.308
High
0311A04
195k / 3.158
1020k / 1.052
Low
0624A01
195k / 13.397
1020k / 3.046
Monitor
0624A05
100k / 8.334
1020k / 3.150
High

## Slide 6
5. Latest 1020k Plate-level Analysis
The latest node includes only plate1-3; all three valid A05 plates are in the 6.7-6.9 range, indicating broad elevation.
Opentrons
P1000H O-ring Data Analysis | 2026-06-04 | 6/13
Latest CV comparison
0
1
2
3
4
5
6
7
8
4.379
0624A04
1.487
1217A05
0.677
0311A04
1.633
0624A01
6.839
0624A05
Unit
Latest CV
Plate1 / Plate2 / Plate3 / Plate4 / Plate5
Valid
Note
0624A04
4.379
4.324 / 4.352 / 4.447 / #DIV/0! / #DIV/0!
3
New optical test standard; only 3 plates tested
1217A05
1.487
0.562 / 0.544 / 0.450 / #DIV/0! / #DIV/0!
3
-
0311A04
0.677
0.549 / 0.571 / 0.577 / #DIV/0! / #DIV/0!
3
-
0624A01
1.633
0.448 / 0.860 / 0.454 / #DIV/0! / #DIV/0!
3
-
0624A05
6.839
6.743 / 6.944 / 6.898 / #DIV/0! / #DIV/0!
3
-
Data-scope risk: the latest 1020k result is not fully comparable with earlier 5-plate data. A05 remains abnormal because all 3 valid plates are elevated, but latest CV comparisons for other units should be interpreted cautiously.

## Slide 7
6. Anomaly Timeline: Relationship Between Aspiration / LOW %D and Optical Peaks
Anomalies do not appear only at the endpoint; key evidence for A04, A01, and 1217A05 is concentrated in the early/mid stages.
Opentrons
P1000H O-ring Data Analysis | 2026-06-04 | 7/13
Cycle
Unit
Observed data / note
Evidence type
100k
0624A04
F1 LOW %D in plate5
Aspiration anomaly
195k
0624A01
H1 LOW %D plate2; G12 LOW %D plate3; worst CV 13.397
Aspiration + channel peak
260k
0624A04
A5 LOW %D plate4; CV 4.214
Aspiration anomaly
455k / 175%
1217A05
Explicit fail; C5/C7 aspirate less
Explicit aspiration fail
520k
0624A04
Oil leak / sensor error; CV 5.642; worst 11.204
Event + optical peak
520k
0624A01
Max CV 8.406
Optical peak
1020k
0624A05
Latest CV 6.839; plate1-3 all high
Endpoint broadly elevated
1020k
All
Only 3 valid plates; plate4/5 #DIV/0!
Data scope change
Analytical focus: O-ring risk is not determined by a single CV value; it is determined by the combined pattern of aspiration anomaly + channel localization + CV peak + event record.

## Slide 8
7. Per-unit Data Matrix: Latest, Peak, and Anomaly Evidence
This page is the unit-level data summary for follow-up review and root-cause investigation.
Opentrons
P1000H O-ring Data Analysis | 2026-06-04 | 8/13
Unit
Risk
Latest CV
Max CV
Max worst CV
Notes / fail evidence
0624A04
High
4.379
520k / 5.642
520k / 11.204
100k F1 LOW %D in plate5; 260k A5 LOW %D plate4; 520k tip-pickup step loss or collision error: oil leak; back board filled with oil, causing sensor error; 1020k new optical test standard, only 3 plates tested
1217A05
High
1.487
520k / 6.195
520k / 7.419
455k C5/C7 two channels aspirated less
0311A04
Low
0.677
0 / 1.957
195k / 3.158
-
0624A01
Monitor
1.633
520k / 8.406
195k / 13.397
195k H1 LOW %D in plate2 and G12 LOW %D in plate3; 920k optical; - optical
0624A05
High
6.839
1020k / 6.839
100k / 8.334
-
High
High
Low
Monitor
High

## Slide 9
8. Case Analysis - 0624A05: Late-stage Optical Spike
The key point for A05 is not historical LOW %D; it is that all three valid plates show high CV at the endpoint.
Opentrons
P1000H O-ring Data Analysis | 2026-06-04 | 9/13
CV trend - all units, A05 highlighted by final spike
0
1
2
3
4
5
6
7
8
9
0
260k
520k
820k
1020k
5
0624A04
1217A05
0311A04
0624A01
0624A05
Cycle
Life
CV
Worst
Note
0
0.0%
0.860
1.212
-
65k
25.0%
1.051
4.609
-
100k
Lifetest Spec
1.300
8.334
-
130k
50.0%
0.641
1.037
-
195k
75.0%
0.915
4.362
-
260k
100.0%
0.602
1.061
-
520k
200.0%
0.748
1.387
-
620k
238.5%
0.656
0.921
-
820k
315.4%
0.547
1.253
-
1020k
392.3%
6.839
3.150
-
Data read: latest CV 6.839 is also max CV, with plate1/2/3 = 6.743 / 6.944 / 6.898. This is a repeatable plate-level pattern at the latest point, not a single plate outlier.

## Slide 10
9. Case Analysis - A04 / 1217A05: Aspiration Anomaly and Local Channel Risk
A04 has LOW %D plus an oil-leak event; 1217A05 has an explicit C5/C7 low-aspiration fail.
Opentrons
P1000H O-ring Data Analysis | 2026-06-04 | 10/13
Unit
Cycle
Data evidence
Interpretation
0624A04
100k
F1 LOW %D in plate5
Early aspiration anomaly
0624A04
260k
A5 LOW %D plate4
Repeated LOW %D
0624A04
520k
Oil leak / sensor error; max CV 5.642; max worst 11.204
Event + optical peak
0624A04
1020k
Latest CV 4.379; only 3 valid plates
Endpoint still elevated
1217A05
455k / 175%
Explicit fail; C5/C7 aspirate less
Explicit aspiration fail
1217A05
520k
Max CV 6.195; max worst 7.419
Optical peak
1217A05
1020k
Latest CV 1.487; worst 2.308
Endpoint improved, but historical fail remains relevant
4.38
A04 latest CV
11.20
A04 max worst
1.49
1217 latest CV
7.42
1217 max worst

## Slide 11
10. Case Analysis - A01 / 0311A04: Recovery Type vs Stable Reference
A01 had large historical variation but improved at the latest node; 0311A04 shows stable optical behavior and can be used as a reference unit.
Opentrons
P1000H O-ring Data Analysis | 2026-06-04 | 11/13
CV trend: A01 recovery vs 0311A04 stable baseline
0
1
2
3
4
5
6
7
8
9
0
260k
520k
820k
1020k
5
0624A04
1217A05
0311A04
0624A01
0624A05
Unit
Cycle
Data
Read
0624A01
65k-520k
CV 4.757 → 8.406 high band
Clear historical variation
0624A01
195k
H1 LOW %D; G12 LOW %D; worst 13.397
Aspiration anomaly + channel peak
0624A01
1020k
Latest CV 1.633; worst 3.046
Improved, but repeatability should be confirmed
0311A04
0-1020k
Latest CV 0.677; max CV 1.957
Overall stable
0311A04
195k
Max worst 3.158
Low-amplitude channel peak; acceptable as a reference
A01 max CV 8.406 vs latest 1.633; 0311A04 latest 0.677. In the data, A01 is a recovery-type unit and 0311A04 is a stable-type unit.

## Slide 12
11. O-ring Evidence Strength Matrix
Separate direct data evidence from O-ring root-cause inference to avoid presenting correlation as causation.
Opentrons
P1000H O-ring Data Analysis | 2026-06-04 | 12/13
Unit
Optical / aspiration evidence
Supporting event
O-ring inference strength
0624A05
Strong: latest CV 6.839; all 3 valid plates are elevated
Weak: no explicit LOW %D note
Medium-high: overall consistency declined; before/after replacement retest is needed
0624A04
Strong: F1/A5 LOW %D; max worst 11.204
Strong: oil leak / sensor event
High: contamination, lubrication, or sealing-surface abnormality requires teardown
1217A05
Strong: C5/C7 low aspiration fail
Medium: 520k CV peak
High: local O-ring or channel assembly risk
0624A01
Medium: H1/G12 LOW %D; historical CV high
Medium: latest node improved
Medium: stage-specific anomaly; repeatability is needed
0311A04
Weak: no anomaly note; CV is stable
Weak
Low: reference sample
Root-cause confirmation standard: teardown finds an O-ring abnormality, and the corresponding channel LOW %D / CV anomaly improves significantly and repeatably after replacement.

## Slide 13
12. Additional Data and Final Judgment Recommendations
The current data is sufficient for risk stratification, but final O-ring performance judgment requires additional verification data.
Opentrons
P1000H O-ring Data Analysis | 2026-06-04 | 13/13
#
Data needed
Details
Purpose
1
O-ring teardown data
Wear, scratches, contamination, lubrication condition, compression set
Connect optical / aspiration anomalies with physical condition
2
Before/after replacement A/B data
Same unit and same channel: CV, LOW %D, and aspiration-volume changes before/after O-ring replacement
Verify whether the O-ring is the main cause
3
Channel-localized retest
Repeat C5/C7, F1/A5, H1/G12, and all A05 channels
Confirm whether the anomaly is fixed to a channel or plate
4
Unify plate-count scope
Complete 5 plates at the latest node or explicitly define the 3-plate standard
Resolve comparability between historical and latest data
5
Final acceptance rule
Combine optical, aspiration, and seal/leak indicators for judgment
Avoid assigning PASS based only on cycles or a single CV metric